## Slide 1
The fairSIM project
Structured Illumination Microscopes and Reconstruction for Video-Rate Imaging
2019-04-15
Marcel Müller1, Andreas Markwirth2, Alice Sandmeyer2, Robin Van den Eynde1, Mario Lachetta2,
Viola Mönkemöller1, Wolfgang Hübner2, Mick Phillips3, Rainer Heintzmann1, Ian Dobbie3,
Lothar Schermelleh3, Peter Dedecker1, Thomas Huser2
1Laboratory for NanoBiology, Department of Chemistry, KU Leuven, Leuven, Belgium.
2Biomolecular Photonics, Faculty of Physics, Bielefeld University, Bielefeld, Germany
3Micron Advanced Bioimaging Unit, Department of Biochemistry, University of Oxford, Oxford, UK.
4Institute of Physical Chemistry and Abbe Center of Photonics, Friedrich-Schiller-University, Jena, Germany

## Slide 2: What kind of ‘SIM’ are we using?
Gustafsson-Heintzmann-style “coherent” SIM
Sinusoidal intensity patterns, with a spatial frequency close to the resolution limit
Generated by interfering (2 or 3) beams of coherent light
Axial resolution enhancement (Gustafsson 2008)
  Three-beam interference pattern
Lateral resolution increase (set by spatial frequency of the pattern)
Nice “push” in medium frequency response
  (see 200nm in the OTF plot)
Direct reconstruction:
  No need for iterative “deconvolution-like” solvers, though they can be applied
Gustafsson (2000) 10.1046/j.1365-2818.2000.00710.x
Gustafsson et al. (2008) 10.1529/biophysj.107.120345
Heintzmann, Cremer (1999) 10.1117/12.336833
Frohn (2000, PhD thesis) 10.1117/12.336833

## Slide 3: What kind of ‘SIM’ are we using?
Gustafsson-Heintzmann-style “coherent” SIM
Sinusoidal intensity patterns, with a spatial frequency close to the resolution limit
Generated by interfering (2 or 3) beams of coherent light
Axial resolution enhancement (Gustafsson 2008)
  Three-beam interference pattern
Lateral resolution increase (set by spatial frequency of the pattern)
Nice “push” in medium frequency response
  (see 200nm in the OTF plot)
Direct reconstruction:
  No need for iterative “deconvolution-like” solvers, though they can be applied
Gustafsson (2000) 10.1046/j.1365-2818.2000.00710.x
Gustafsson et al. (2008) 10.1529/biophysj.107.120345
Heintzmann, Cremer (1999) 10.1117/12.336833
Frohn (2000, PhD thesis) 10.1117/12.336833
Simulated intensity profile for 2-bam and 3-beam SIM, x/z. See also
Gustafsson 2008 for experimental data.

## Slide 4: What kind of ‘SIM’ are we using?
Gustafsson-Heintzmann-style “coherent” SIM
Sinusoidal intensity patterns, with a spatial frequency close to the resolution limit
Generated by interfering (2 or 3) beams of coherent light
Axial resolution enhancement (Gustafsson 2008)
  Three-beam interference pattern
Lateral resolution increase (set by spatial frequency of the pattern)
Nice “push” in medium frequency response
  (see 200nm in the OTF plot)
Direct reconstruction:
  No need for iterative “deconvolution-like” solvers, though they can be applied
Gustafsson (2000) 10.1046/j.1365-2818.2000.00710.x
Gustafsson et al. (2008) 10.1529/biophysj.107.120345
Heintzmann, Cremer (1999) 10.1117/12.336833
Frohn (2000, PhD thesis) 10.1117/12.336833
SIM vs. Confocal vs. Widefield

## Slide 5: What kind of ‘SIM’ are we using?
Gustafsson-Heintzmann-style “coherent” SIM
Sinusoidal intensity patterns, with a spatial frequency close to the resolution limit
Generated by interfering (2 or 3) beams of coherent light
Axial resolution enhancement (Gustafsson 2008)
  Three-beam interference pattern
Lateral resolution increase (set by spatial frequency of the pattern)
Nice “push” in medium frequency response
  (see 200nm in the OTF plot)
Direct reconstruction:
  No need for iterative “deconvolution-like” solvers, though they can be applied
Gustafsson (2000) 10.1046/j.1365-2818.2000.00710.x
Gustafsson et al. (2008) 10.1529/biophysj.107.120345
Heintzmann, Cremer (1999) 10.1117/12.336833
Frohn (2000, PhD thesis) 10.1117/12.336833
SIM vs. Confocal vs. Widefield
SIM excels at
Straight-forward resolution doubling on “nice” samples
Lots of speed for live-cell imaging
This talk:
Software for reconstruction
Microscopes
Integrated video-rate SIM

## Slide 6: Software

## Slide 7: How the fairSIM project came to be
Early 2014 (five years ago...)
SIM algorithm published for many years, but no freely available implementation
Building SLM-based SIMs started to get popular
Localization community: Lots of open software solutions (QuickPALM, rapidSTORM).
Early 2016 – fairSIM “1.0”
Handles 2-beam and 3-beam data
Runs as an ImageJ / FIJI plugin
(hopefully) user-friendly
Public, open-source, published www.fairsim.org
Also: SIMToolbox and OpenSIM
Müller et al. (2016) 10.1038/ncomms10980
fairsim.org
Pavel Křížek et al. (2016) 10.1093/bioinformatics/btv576
Amit Lal et al. (2016) 10.1109/JSTQE.2016.2521542

## Slide 8: How the fairSIM project came to be
Müller et al. (2016) 10.1038/ncomms10980
fairsim.org
Pavel Křížek et al. (2016) 10.1093/bioinformatics/btv576
Amit Lal et al. (2016) 10.1109/JSTQE.2016.2521542

## Slide 9: Software, example data, how-to’s
Single slice SIM reconstructions from fairSIM (data taken on a DeltaVision|OMX v4, excitation at 642nm):
200nm fluorescent beads, wide-field (a) and SIM (b).
CellMask Deep Red membrane stain on LSEC, wide-field (c) and SIM (d)
Scalebar 5µm
Müller et al. (2016) 10.1038/ncomms10980
fairsim.org

## Slide 10: Single-slice SIM: OTF attenuation / notch filtering
Out-of-focus light („background“) leads to artifacts in the reconstruction
In terms of OTFs: Missing cone problem in the 3D OTF
Can be mitigated in 2D SIM (via 3-beam or lower lateral resolution)
Almost required for non-TIRF datasets
Freely configurable in fairSIM
K. Wicker, R. Heintzmann, Opt. Exp. 2013, doi:10.1364/OE.21.002032
K. O’ Jolleran, M. Shaw, Biomed. Opt. Exp. 2014, doi:10.1364/BOE.5.002580
M. Shaw, K. O’ Jolleran, Optics Letters, 2012, doi: 10.1364/OL.37.004603
from K. Wicker, Opt. Exp. 2013, fig. 7 (open access)

## Slide 11: Single-slice SIM: OTF attenuation / notch filtering
Actin filaments in U2OS cells, labelled with Phalloidin-Atto488, measured on
a DeltaVision|OMX, excited at 488 nm wavelength.
reconstruction using a standard 2D OTF
reconstruction using the attenuated OTF (strength 0.995, FWHM 2 cycles/micron).
Scalebar 5µm, inset 2µm
Müller et al. (2016) 10.1038/ncomms10980
fairsim.org

## Slide 12: Project in active development
Time-lapse support (v.1.1.0-beta2)
Subscribe to updates:
github.com/fairSIM
@mueller_physics on Twitter

## Slide 13: Project in active development
Image corrections and parameter feedback (v.1.3.0)
Subscribe to updates:
github.com/fairSIM
@mueller_physics on Twitter

## Slide 14: Project in active development
Subscribe to updates:
github.com/fairSIM
@mueller_physics on Twitter
Connectivity to
Java
Beanshell
(other ImageJ languages)
Matlab
Inspired by a collaboration in Leuven running high-throughput imaging
Scripting support (v.1.4.2)

## Slide 15: fairSIM – full 3D
Summer project at Micron Oxford: Full 3D reconstruction in fairSIM
‘develop-3D-SIM’ branch (github)
Working beta version, currently fixing bugs
Actin filaments in LSEC cells, labelled with Phalloidin-Atto488, measured on a DeltaVision|OMX, excited at 488 nm. Scalebar 5µm, single slice vs. 2µm z-projection

## Slide 16: “fairSIM 3D” – new GUI, new concept
3D SIM processing requires more knowledge about the microscope (esp. OTFs)
Editor for alignment data (stored in one XML-file)
GUI for multi-channel, multi-timepoint reconstruction
Both parts mostly work, but are very “unpolished”
Known bugs

## Slide 17: fairSIM 3D – OTF generation
OTF phase unwrapping
Single bead, multi-plane stack, under SIM illumination (see Demmerle et. al. for protocol)
Band-separate
Compensate bead position (phase)
Compensate artifacts
Compensate noise / average / smooth
Have a (at least minimalistic) GUI
Test and integrate

## Slide 18: Microscopes

## Slide 19: The ‘fastSIM’ approach
2-beam SIM based on a spatial light modulator
Established concept, used with some variations by many groups: Jena, Oxford, Cambridge, Bielefeld, Leuven
Janelia Farm: Dong Li et. al., Science 2015 on non-linear SIM (>150 pages of supplementals!)
Tweaked for speed: > 60 fps reconstructed
The hardware can provide the speed, the dyes not so much
Hui-Wen Lu-Walther et al. (2015) 10.1088/2050-6120/3/1/014001
Dong Li et al. (2015) 10.1126/science.aab3500
Laurence J. Young et al. (2016) 10.3791/53988

## Slide 20: Why (not) build a SIM
Pro’s:
Easy customization (optics, mechanics, electronics, software)
Integration in existing imaging systems
Specific optimization (speed, uncommon wavelength, …)
Con’s:
Needs some expertise in optics (harder than widefield, easier than STED, comes with its own quirks)
Needs time & dedicated person
Probably won’t reach the stability & uptime of a commercially available solution
Main reason to build a SIM:
tailoring a system to a specific use case
R. Van den Eynde

## Slide 21: Integrating microscope & reconstruction
The fairSIM project
Multi-color extension of the ‘fastSIM’
  488, 568 and 647nm excitation
  Three cameras for fast readout
  Timing control electronics
GUI for system control
  Connects to hardware components, µManager, bUnwarpJ, etc.
A. Markwirth
M. Lachetta

## Slide 22: Integrating microscope & reconstruction
The fairSIM project
On-the-fly reconstruction
  GPU-based
  Also speeds up regular fairSIM
All code on github: fairSIM-VIGOR
Manuscript in review

## Slide 23: Thanks & Questions
Open-source SIM reconstruction software
Video-rate 2D SIM with immediate reconstruction
Compact / mobile microscopes
fairsim.org
mueller-physics.org
Implementing 3D reconstruction
L. Schermelleh
I. Dobbie
M. Phillips
P. Dedecker
R. Van den Eynde
V. Mönkemöller
Customizing SLM-based SIM
T. Huser
A. Markwirth
M. Lachetta
A. Sandmeyer
W. Hübner
R. Heintzmann